

BIR shielding calculation tool

Requirements:

* Exi log from device
  + Must be of a type recognised by the script currently this includes
    - Siemens Ysio
* Room floor plan
* Software:
  + ImageJ/FIJI (FIJI is better for large floor plan files)
  + Local python installation or NAS access
    - For local installation, the packages required are: pandas, numpy, matplotlib, shapely, geos, descartes, pdf2image, glob
  + Set of python scripts
    - Y:\BTS-Herston-MedPhys\MedPhys\BTS Central\Other\Shielding\parse\_exi
    - <https://github.com/ckswilliams/parse_exi>

**Procedure**

1. Make a copy of the input\_sources.csv, input\_rooms.csv, and put them in a directory with the exi\_log (not strictly required, but grouping the files is convenient)
2. Make the floor plan available in an image format
   1. The script pdftoimg can be used for this purpose, it breaks apart a pdf file into an image corresponding to each page
3. Open the image in ImageJ
4. Set the scale of the image in ImageJ
   1. Use the line selector to select a feature of known size
   2. Open Analyse -> Set scale
   3. Input the size of the known feature, in m
   4. Activate the Global checkbox
5. If the floorplan contains multiple x-ray rooms, select the region around one and duplicate it (ctrl–d)
6. Open Analyse -> Set measurements. Change the settings so that only the Bounding rectangle and inverted Y coordinate checkboxes are selected
7. Open the copy of ‘input\_rooms.csv’. For each radiation sensitive room around the x-ray room in the floor plan:
   1. Record the name of the room in the input\_rooms.csv file NO REPEATED ROOM NAMES, DO NOT INPUT THE X-RAY ROOM ITSELF
   2. Select the rectangle tool. For each nearby room, click and drag the rectangle tool to cover the internal wall.
   3. Save the measurement with ctrl-m
   4. Repeat for each radiation sensitive room
8. Copy BX, BY, Width, Height into the similar columns in the input\_rooms.csv file
9. Open the ‘input\_sources.csv’ text file
   1. Either use the multipoint tool or cursor to determine the location of each source and target position for the table, wall bucky, cross shot, 2nd cross shot and 3rd cross shot. An example of these positions is shown here, though professional judgement will be required
   2. Input these values into the ‘input\_sources.csv’ text file
10. To allow the script to create a pretty overlaid image, save the room image in .tif format (probably into the same directory as the input files in order to keep things tidy)
11. Open the script gui using either the run\_interface - network.bat file, or using your local copy of python
12. Select the exi file, input\_sources and input\_rooms files using the menus
13. Click load\_input\_files
    1. If the loading failed, make sure the files were properly selected
    2. Ensure the exi log is from a type of device mentioned at the beginning of this document
    3. Ensure the column headings and data types of the input files are similar to those in the sample files
    4. If all these fail, contact me and I will attempt to troubleshoot
14. Use the show room, Workload, and OGP stats buttons to generate plots in order to sanity check the input files. Use best judgement to evaluate whether the resulting plots appear reasonable
15. Use Lead requirement button to calculate the lead requirement for each room
16. Save plots in order to produce useful reporting diagrams and tables
17. Save verbose dose calculations allow the user to sanity check the calculations and intermediate steps, providing a range of geometric calculations and dose calculations in tabulated files
18. Use the files and results created using this program to inform the x-ray shielding report
19. Repeat ad nauseum and suggest any missing features to Chris

Optional XRAYBARR workflow:

Optionally, ‘Create XRAYBARR spectrums’ allows the user to generate files for use with the XRAYBARR program, as an alternative dose calculation method. These calculations can be performed in a steamlined way using the geometrical data and workloads created.

1. Create input files and save exi logs as per instructions above
2. Perform a full export as per above instructions
3. Open XRAYBARR
4. Select ‘Open Barrier Setup File’
5. Navigate to the designated output directory
6. Decide which wall to calculate the shielding requirement for
7. Open the .bar file for the room likely to correspond to the highest shielding requirement for the chosen wall
8. Open each of the spectrum files corresponding to the different source positions
9. At this point, note the following:
   1. Pats/wk and Tot W mAmin/week will show the wrong values. The correct values, which can be seen in ‘Plot distribution’ are taken from the exi logs. Reloading the spectrum file will fix the display if necessary.
   2. The leakage technique factors are conservative
   3. The primary use factor is binary, based on the same assumptions made during analysis of the exi files
   4. The Grid&Cassette attenuation of the primary beam option is selected
   5. It would be very prudent to sanity check all distances, use factors, and constraints.
10. Allow XRAYBARR to perform the calculation
    1. Enjoy not having to manually input all the distances

Technical specifications:

References:

NCRP 49

NCRP 147

BIR

Archer